Training Notes TcpO2 and TBI training: 			15th and 16th March 2017

John Juter (Perimed) provided the training in house to Vascular testing staff over 2 days. I was in full time attendance over the 2 days due to being the lead in the service development.

# Transcutaneous oximetry TcpO2

# Transcutaneous oximetry, tcpO2, is a non-invasive measurement reflecting the amount of O2 that has diffused from the capillaries through the epidermis.Top of Form

TcpO2 provides continuous information about the body’s ability to deliver oxygen to the tissue. TcpO2 is dependent on oxygen uptake in the respiratory system, the oxygen transport/capacity of the blood and the general status of the circulatory system. Any impairment of the organism's ability to deliver oxygen to the tissue will be revealed immediately since the skin is ranked very low in the body's system of oxygenation priority.

Technically, the electrode heats the underlying tissue to create a local hyperaemia, which intensifies the blood perfusion, increasing the oxygen pressure. In addition, the heat will dissolve the lipid structure of the dead, keratinized cells in the epidermal layer making the skin permeable to gas diffusion. On its way the oxygen may be consumed by the cells if the metabolism is high.

Clinical applications

- Determination of peripheral vascular oxygenation

- Quantification of the degree of peripheral vascular disease

- Determination of the optimum level of amputation

- Evaluation of revascularization procedures

TcpO2 measurements usually require at least two or three sites to provide a good picture. The more sites assessed, the better the oxygenation picture.Bottom of Form

Reference values

50-70 mmHg Normal

< 40 mmHg Impaired wound healing

< 30 mmHg Critical Limb Ischemia

Notes regarding the Perimed machine

The machine has 2 different types of sensor, 2 laser Doppler sensors for performing toe pressures and three tcpO2 sensors, both these sensors have different maintenance requirements.

Both types of sensor cable need to be tucked away neatly in the side baskets, as the drawer is easily closed onto the cables.

Laser Doppler sensors

The laser Doppler sensors are attached to the articulated arm with the cuff inflators, so are easily distinguished.

The laser Doppler sensors are attached via optical cables, these are more rigid and brittle than a standard cable, so we need to be careful not to over bend them as they will fracture.

Cleaning of the sensor and cables can be done using the green Clinell wipes.

TcpO2 sensors

The TcpO2 sensors are colour coded and attach to the front of the Perimed machine for storage whilst the machine is in use. They are currently set to automatically calibrate when the machine is switched on.

The tcpO2 sensors should display an inactive calibrated reading on the machine of around 155-160. When the difference between any two sensor readings exceeds around 15, the membrane on the TcpO2 sensor should be replaced.  Alternatively the sensor with the outlying reading should not be used for any studies until replaced. The box of membranes is expensive at ~£300 per box, so before replacing please see Ellie to go through the technique for how to - as this was tricky.

The tcpO2 sensors should be stored in the supplied caps with a small quantity of saline when they are not going to be used for a few days – this will maintain the membrane and help each one last for longer. There are spare caps provided in the new membrane applicator boxes.

Cleaning of the tcpO2 membrane:  we can use small alcohol wipes.  John wasn’t able to confirm that use of clinell wipes would be okay and not damage the delicate membrane.

Blood pressure cuffs

A variety of blood pressure cuffs have been supplied. Blue arm and ankle cuffs, which are colour coded to the machine schematic for ease of cuff placement. Small white digit cuffs of varying sizes for digits. Standard cleaning and cuff maintenance applies. Standard Clinell wiping or use soap and water when visibly soiled.

The small toe pressure cuffs should last up to a year (depending on usage) and should never be inflated when not attached properly to a digit as they will pop. When on the toe/finger there should always be an overlap of the bladder around the finger otherwise they may become damaged.

Using the Machine

We had a practice session with a few patients in the afternoon on day 1 and 2. Patient details were manually entered onto the system with the hospital number entered as patient ID. We attached the laser Doppler probes and selected the correct cuffs for toe and arm pressures; then connected the pneumatic hoses to the correct cuffs (colour coded).  Once the TBI had been performed the software then selected the tcpO2 setting for 15 minutes of monitoring, averaging to the final value over a stable (plateaued) 30 second time interval.

The training provided a good foundation for carrying on and developing this new service within North Bristol NHS Trust.