Public Health England; Infection Prevention and Control, Use of Ultrasound Gel

Webinar 18.03.2021

The MHRA are undergoing investigation in the three clusters of Bcc (Burkholderia Cepacia Complex) outbreak where the cases are linked as having had ultrasound scans. One long-running cluster identified Bcc isolated from ultrasound gel samples submitted from 8 Trusts that contained the same strain of Bcc as from the patients.

The conclusion is that there is a risk of cross infection to immunosuppressed patients through the use of ultrasound gel.

PHE have published guidelines for optimum practice in IP&C regarding the supply, storage and use of ultrasound gel. These guidelines have been circulated throughout our department.

Following the presentation there was a Q&A Session: A panel of experts answered questions posted by delegates during the presentation.

Warming the gel – opinions?

Increasing the temperature to body temperature will increase the replication of pathogens. Gel warming is NOT A GOOD IDEA

Immuno-suppressed patients – what’s the definition?

This is difficult to define. More vulnerable to infection? Contraindicated for a live vaccine?

In practice with ultrasound the relevant issues are skin breaks and mucous membranes. The clusters under investigation all had multiple skin breaks or devices that crossed mucous membranes.

The important considerations are BROKEN, EXPOSED SKIN and the IMMUNOCOMPROMISED

Cost implications / Environmental implications?

NHS supply chain has been in contact with suppliers. Currently two products have been suspended. The recommendations are based on optimum safety but are GUIDANCE

Could other products contribute?

Good disinfectant is now very safe. Contamination from probe disinfectant is very rare

Should there be a call for single use containers only or for decanting from large containers stopped?

Yes – in an ideal world, but  there are currently financial implications and we need to be pragmatic. Ask yourself – if you were the patient what gel would you like used?

There is a growing movement away from 5litre gel containers.

How should we handle the bottle nozzle?

Carefully, with clean hands! Clean all external spills on the bottle. Be very careful opening the nozzle

How to store bottles after opening?

There is a huge gap in the guidance market around ‘gel’ in general. This needs addressing.

What about lubricating gel?

Sterile single use sachets only!