Test Protocol: Bypass Graft Ultrasound Scan

Protocol last updated on:  April 2018

Graft surveillance is carried out on all vein grafts and selected prosthetic grafts at the clinicians request. Vein grafts are scanned along their entire length to identify atheroma and evaluate any stenotic lesions affecting the graft or anastomoses. The inflow and run-off are also assessed.

Early detection and treatment of lesions (by angioplasty, thrombolysis or surgery) can restore blood flow to the foot and muscle compartments and avert graft failure.

Equipment required:

High quality Duplex ultrasound machine:

Cart based ultrasound system e.g. Philips Epiq5; including low, mid, and high frequency probes (ranging 1MHz-17Mhz)

Portable ultrasound system e.g. Philips CX50; including low, mid, and high frequency probes (ranging 1MHz-17Mhz)

Peripherals:

CRIS, PACS

Consumables:

Paper roll, paper hand towels, ultrasound gel, Clinell wipes

Other:

An adjustable examination couch is required.  If the examination is performed with the operator seated, a height adjustable chair with good lumbar support is required. The examination room should be temperature controlled with adjustable lighting levels suitable for the examination.

Quality Analysis & Calibration:

Imaging: Serviced and safety tested every 12 months. Annual quality assurance testing by qualified personnel from radiation protection service (RPS). Review of in-service equipment should typically be undertaken four to six years after installation.
Doppler: none

Patient and staff safety:

Electrical: Equipment must comply with the Medical Devices Directive. Manufacturer completes PPQ (post-purchase questionnaire) form. Safety tested on arrival and then annually.

Infection control:

Appropriate levels of PPE should be worn for the examination (barrier guidelines should be adhered to for infectious patients). Probes should be thoroughly cleaned with a Clinell disinfectant wipe between all patients.

In cases where infectious patients have been assessed (e.g. MRSA and Clostridium difficile) the ultrasound machine should be thoroughly cleaned with Clinell wipes post-examination.

For sterile procedures sterile probe covers and gel should be used.

Ultrasound Safety:  Output powers are quoted by manufacturer.

Users are required to be aware of the potential risks and methods of minimising ultrasound exposure (guide-lines: The safe use of Ultrasound in Medical Diagnosis. British Medical Ultrasound Society. Ed. G. ter Haar and F A Duck. 2000).

Communication with patients:   The vascular scientist undertaking the examination should:

- Introduce themselves

- Confirm the patients identity e.g. full name and date of birth

- Explain why the examination is being performed

- Give an explanation of the procedure and its duration and the reason for the scan – consideration should be made to the age and mental status of the patient

- Obtain verbal consent for the examination

- Obtain a pertinent medical history from the patient and/or notes:

  - Verify that the requested procedure correlates with the patient’s clinical situation

  - Type of and position of graft e.g. fem-distal to PTA, fem-pop

  - History of previous treatment e.g. angioplasty

  - Results of other relevant diagnostics & previous vascular studies

  - Any recurrence of symptoms or associated problems with the leg

- Explain how/when the patient will be informed of the results

Test procedure

Patients attend for graft surveillance at week 6 (in consultant clinic); then at months 3, 6, 9, 12 18 and 24 for vein grafts (in surveillance clinics).  Patients will be discharged from graft surveillance after 2 years providing there are no issues with the graft.

Femoro-popliteal and fem-distal grafts:

Upper anastomosis:

Using the 12-3 MHz transducer and with the patient supine, the test is commenced at the level of the groin where an assessment can be made of the common femoral artery (CFA) waveform. Demonstration of a sharp upstroke and a biphasic/triphasic signal rules out the necessity to scan the iliac vessels.  Any deviation from the normal waveform and flow indicates iliac vessels should also be included in the scan.

Proximal anastomosis must be visualised in B mode and examined using colour and spectral Doppler to confirm normal biphasic/triphasic flow.  Any atheroma or stenosis must be noted.  The graft must be followed using colour Doppler to check for patency, dilatations, kinking, atheroma, and stenosis, and a normal waveform demonstrated in the mid graft.  Distal anastomosis must also be examined as above for proximal anastomosis.  Distal run off vessel must be examined using colour and spectral Doppler to confirm biphasic/triphasic waveform and normal flow velocities.

If the graft is to:

a) Above and below knee popliteal/TP trunk it is necessary to scan the popliteal artery and at least one tibial vessel to the ankle (see lower limb arterial protocol).  This should be the best/dominant vessel, which can usually be identified from previous scans.

b) One of the tibial arteries (AT, PT or peroneal) - it is necessary to scan the vessel to the ankle.

N.B. Be aware that it is possible to see flow disturbance at the anastomoses sites due to non-physiological geometry of anastomoses e.g. a difference in diameter between native vessel and graft.

Iliofemoral and fem-fem cross-over grafts:

A full aorto-iliac scan should be carried out on the donor side (as outlined above). Both anastomoses should be assessed in detail and the presence of patent SFA or profunda origins confirmed or refuted on the recipient side.

Interpretation and Grading of disease

Graft and vessels are scanned to determine patency and identify stenoses, atheroma, and dilatations.

The extent of atheroma found is assessed by estimating any reduction in the vessel lumen apparent from the image. Filling defects and colour flow disturbances serve to alert the operator to potential sites of stenoses. Colour flow changes will be most apparent if a PRF is selected which places normal flow velocities in the upper part of the colour scale. Higher velocities will then appear as areas of aliasing.

Percentage stenosis is finally determined by comparing peak systolic velocity proximal to a stenotic area with intra-stenotic peak velocity. Thus a peak systolic velocity ratio (PSV ratio) is produced which correlates to a particular degree of narrowing (see table).

The presence of hyperaemia, high resistance signals and particularly low flow should be noted.

N.B. In the single visit clinic setting limited investigations may be carried out as per instruction from the requesting clinician.

Reporting of results:

The report is should be written by the vascular scientist undertaking the examination.

The report should include the correct patient demographics, date of examination and examination type. The report should include:

- Location of the graft, i.e. fem-pop, fem-peroneal etc

- Quality of the inflow to the graft

- State of proximal and distal anastomoses

- Details of any stenosis, atheroma or dilatations, kinks in the graft and any native vessel under investigation

- Comments on patency and quality of flow within the graft and the connected run-off vessel should be made

- Waveforms should be included to support the report

- Velocity increases are indicated in a site-specific fashion.

If no problems have been detected in the graft, patients are advised that all is well and given an appropriate routine follow up appointment.

If moderate disease is noted (i.e. 2-3x velocity increase) the patient is advised that they should return for earlier surveillance. The relevant consultant is informed by email.

If a >3x velocity increase is found, the vascular scientist may tactfully indicate that immediate advice will be sought from a vascular surgeon. In this case a provisional appointment may be given, or may be delayed until advised accordingly. In most cases, lesions above 3x are deemed suitable for intervention.

In the case of very severe disease or impending graft failure, the patient should be asked to wait until a vascular surgeon has been consulted. Graft patients can be discharged after a period of 2 years with no significant graft problems.

N.B. In the single visit clinic setting limited investigations may be carried out as per instruction and this shall be acknowledged in the report.

| Anastomosis PSV m/s | Action |  |
| --- | --- | --- |
| <1.5 m/s | Normal Continue surveillance |  |
| >1.5 and <2.8 m/s | Mid graft flow <0.5 m/s | Intervene |
|  | Mid graft flow >0.5 m/s | Inform consultant Continue surveillance |
| >2.8 m/s | Intervene |  |

| In graft stenosis (PSR) | Action |  |
| --- | --- | --- |
| < 1.8 | Normal Continue surveillance |  |
| 1.8 – 2.8 | Mid graft flow <0.5 m/s | Intervene |
|  | Mid graft flow >0.5 m/s | Notify consultant Continue surveillance |
| >2.8 | Intervene |  |
| Occlusion | Inform consultant Discontinue surveillance |  |

| Mid graft velocities | Action |
| --- | --- |
| 1 – 1.5 m/s | Normal Continue surveillance |
| 0.5 – 1 m/s | Inform consultant of reduced velocities within the graft Continue surveillance |
| >1.5 m/s | Inform consultant of increased velocities within the graft Continue surveillance |
| < 0.5 m/s | Inform consultant of potential graft failure Continue surveillance |

References:

Brennan. J. et al (1991). The role of simple non-invasive testing in infra-inguinal vein graft surveillance. European Journal of Vascular Surgery. 5: 13-17.

Idu. M. et al. (1999). ‘Factors influencing the development of vein-graft stenosis and their significance in clinical management.’ European Journal of Vascular and Endovascular Surgery. 17: 15-21.

London. N. (1993). interventional radiology in the maintenance of infrainguinal graft patency.’ British journal of Surgery. 80: 187-193.

Polak. J. et al (1990). ‘Early detection of saphenous vein arterial bypass graft stenosis by colour-assisted duplex sonography.’ American Journal of Radiologx. 154: April.

Savers. R. et al (1993). ‘The histopathology of infrainguinal vein grafts.’ European Journal of Vascular Surgery. 7: 16-20.

Thrush. A. (2000). ‘The role of ultrasound in lower limb arterial bypass graft surveillance.’ BMUS Bulletin. August. 10-13.

Zeibel. W. (1992). Introduction to Vascular Ultrasonography. Philadelphia. W. B. Saunders Company.