Test Protocol: Lower Limb Arterial Ultrasound Scan

Protocol last updated: April 2018

Arterial duplex ultrasound is performed on patients with a history of, or suspected, intermittent claudication, critical limb ischaemia or tissue loss to ascertain the presence and site of stenotic or occlusive lesions affecting the lower limb arteries. These lesions can reduce blood flow to the foot and muscle compartments of the limb.

Duplex scanning allows imaging of the lower limb vessels from the aorta to the ankle. The colour Doppler facility enables patency of the vessels to be determined and highlights areas of velocity increase. The spectral Doppler facility is used to estimate percentage stenosis and to determine a range of flow-limiting or flow-enhancing states.

Equipment required:

High quality colour Duplex ultrasound machine:

Cart based ultrasound system e.g. Philips Epiq 5G; including low, mid, and high frequency probes

Portable ultrasound system e.g. Philips CX50; including low and mid frequency probes

Peripherals:

CRIS, PACS

Other:

Electronic adjustable examination couch, height adjustable chair

Consumables:

Paper roll, paper hand towels, ultrasound gel, clinell wipes

Quality Analysis & Calibration:

Imaging: Serviced and safety tested every 12 months.  Annual quality assurance testing by Radiation Protection Services (RPS).  
Doppler: none

Patient and staff safety:

Electrical:

Equipment must comply with the Medical Devices Directive. Manufacturer completes PPQ (post-purchase questionnaire) form. Safety tested on arrival.

Infection control:

Appropriate levels of PPE should be worn for the examination (barrier guidelines should be adhered to for infectious patients). Probes should be thoroughly cleaned with a Clinell disinfectant wipe between all patients.

In cases where infectious patients have been assessed (e.g. MRSA and Clostridium difficile) the ultrasound machine should be thoroughly cleaned with Clinell wipes post-examination.

For sterile procedures sterile probe covers and gel should be used.

Ultrasound Safety:

Output powers are quoted by manufacturer.

Users are required to be aware of the potential risks and methods of minimising ultrasound exposure (guidelines: The safe use of Ultrasound in Medical Diagnosis. British Medical Ultrasound Society. Ed. G. ter Haar and F A Duck. 2000).

Communication with patients:   The vascular scientist undertaking the examination should:

- Introduce themselves

- Confirm the patients identity e.g. full name and date of birth

- Explain why the examination is being performed

- Give an explanation of the procedure and its duration – consideration should be made to the age and mental status of the patient

- Obtain verbal consent for the examination

- Obtain a pertinent medical history from the patient and/or notes:

  - Verify that the requested procedure correlates with the patient’s clinical situation

  - Results of other relevant diagnostics & previous vascular studies

- Explain how/when the patient will be informed of the results

Test procedure

The patient must be capable of lying supine and still during the scan. Where appropriate, an ability to lay flat assists greatly in visualisation of the aorto-iliac segment.

Femoral-popliteal segment:

Using the 12-3MHz transducer and with the patient supine, the test is commenced at the level of the groin where an assessment can be made of the common femoral arterv (CFA) patency and waveform. Demonstration of a sharp upstroke and a biphasic/triphasic signal usually rules out the probable presence of iliac disease. The profunda femoris artery (PFA) and superficial femoral artery (SFA) origins are identified and assessed for patency and stenoses. The SFA is then followed throughout the length of the thigh using an anterior-medial window and with some knee flexion where helpful. In the lower thigh, the SFA descends deep into the adductor canal before re-emerging as the popliteal artery behind the knee. The popliteal artery is best visualised by lying the patient on their side and should be assessed throughout its length. Spectral Doppler should be used to obtain waveforms in the following locations: CFA, PFA origin, SFA prox, mid, distal, popliteal mid and distal and TPT (if visualised). In the presence of stenosis, waveforms should be taken where appropriate to quantify severity. If there is an occlusion, images should be taken to demonstrate absence of flow, location of the occlusion and where flow restores.

Crural vessels:

The ATA, PTA and peroneal artery signals are assessed at ankle level using spectral Doppler. A comparison is made to the Doppler signal obtained from the distal popliteal artery in order to assess the likelihood of crural vessel disease.

Aorto-iliac segment:

If monophasic signals are detected in the CFA, or if specifically requested by the referring doctor, the aorto-iliac segment is also assessed. Colour Doppler is used to determine the patency of the distal aorta, common iliac artery (CIA) and external iliac artery (EIA). Spectral Doppler is used to assess the flow in these vessels. Additional spectral Doppler traces should be obtained in the presence of stenosis to determine the severity. If an occlusion is detected, colour Doppler imaging should be used to demonstrate the absence of flow and determine where it restores.

# Interpretation and Grading of disease

The peripheral arteries should be scanned in an attempt to determine patency and identify the presence of disease (i.e. diffuse disease, stenoses, occlusion, and aneurysm) with a view to producing a site-specific, representative map of the peripheral arterial circulation.

Percentage stenosis is determined by comparing peak systolic velocity proximal to a stenotic area with intra-stenotic peak velocity. Thus a peak systolic velocity ratio (PSV ratio) is produced which correlates to a particular degree of narrowing.

| Diameter Reduction | Velocity Ratio (Vs/Vp) |
| --- | --- |
| < 50% | <2 |
| > 50% | >2 |
| >75% | >4 |
| Occluded | No flow detected |

N.B. In the one-stop clinic setting limited investigations may be carried out as per instruction from the requesting clinician.

Reporting of results:

The report is written on CRIS by the vascular scientist who performed the examination.

The report should include the correct patient demographics; date of examination and examination type.

The report should include:

- Presence of any diffuse disease/atheroma/calcification

- Patency and flow pattern/waveforms seen within the CFA, PFA origin, SFA and popliteal artery

- Distal run off (patency of distal vessels, waveforms)

- Presence, severity and location of any stenoses or occlusions

- Any incidental findings

- Limitations of the examination e.g. acoustic shadowing due to calcification

The following optimised set of images should be obtained and stored:

- CFA colour and spectral Doppler trace with PSV measurement

- SFA (proximal, mid and distal) colour and spectral Doppler traces with PSV measurements

- Popliteal artery (mid and distal) colour and spectral Doppler traces with PSV measurements

- Colour Doppler image of the tibio-peroneal trunk (if visualised)

- Distal PTA, ATA and peroneal artery colour spectral traces with PSV measurements

In the presence of disease, additional images should be taken to demonstrate disease as required.

References:

Allard. L. et al,  (1999). ‘Review of the assessment of single level and multilevel arterial occlusive disease in lower limbs by duplex ultrasound’. Ultrasound in Medicine and Biologv. 
25(4): 495-502.

Hatsukhami. T. et al. (1992). Colour Doppler Characteristics in Normal Lower Extremity Arteries. Ultrasound in Medicine and Biology. 18(2): 167-171.

Lunt. Ni. (1999). Review of duplex and colour Doppler imaging of lower-limb arteries and veins. Journal of Tissue Viability. 9(2): 45-55.

Sensier. Y. Bell. P.. London. N. (1998). ‘The ability of qualitative assessment of the common femoral Doppler waveform to screen for significant aortoiliac disease’. European Journal of Vascular and Endovascular Surgery. 15(4): 35 7-64.

Sensier. Y. et al. (1996). ‘A Prospective Comparison of Lower Limb Colour-coded Duplex 
Scanning with Arteriographv’. European Journal of Vascular and Endovascular Surgerv. 11: 
170-175.